## Slide 1: Water Chemistry Analysis
Sampling
  Where?
  When?
  How Big?
  Filter?
  Preserve?

## Slide 2: Analysis Techniques
Electrochemical
Spectroscopic
Chromatographic
How do these ‘probe’ different properties?

## Slide 3: Electrochemistry
Common techniques: pH, conductivity, Ion-specific electrodes, amperometry, voltammetry
Utilize electrical properties of analyte to determine concentrations

## Slide 4: Electrodes
The parts of a cell that are the sites of anodic and cathodic reactions
Working and counter electrodes - conductive materials where the redox reactions occur
Most cells also use a reference electrode  recall that all Energy is relative to a predetermined value, the reference electrode provides an ‘anchor’ for the system
The electrical measuring device measures a property of the electrodes in a solution

## Slide 5
Pathway of a general electrode reaction
electrode
bulk
Interface - electrode surface region
Notes: Critical to review rates of these processes…
Check nonpolarized-polarized diagram – difference is current measure – more current (faradaic or nonfaradaic) results in polarization

## Slide 6
Electroanalytical methods
Voltammetry I=f(E)
bulk methods
Conductometry G=1/R
Conductometric titrations
Interfacial methods
Static methods i=0
dynamic methods i > 0
potentiometry E
Potentiometric titrations volume
Constant current
Controlled potential
Electrogravimetry (wt)
Coulometric titrations Q=it
Electrogravimetry (wt)
Amperometric titrations volume
Constant electrode potential coulometry Q =  i dt

## Slide 7: Potentiometry
Measurement of potential in absence of applied currents
Depends on galvanic cells – no forced reactions by applying external potential
Ecell=Eindicator-Ereference+Ejunction
Reference electrodes are a potentiometric application that stays at a defined (Nernst), constant, and steady potential
  Calomel  Hg/Hg2Cl2(sat’d),KCl(x M)║
  Ag/AgCl  Ag/AgCl(sat’d), KCl (x M) ║

## Slide 8
pH = - log {H+}; glass membrane electrode
pH electrode has different H+ activity than the solution
SCE // {H+}= a1 / glass membrane/ {H+}= a2, [Cl-] = 0.1 M, AgCl (sat’d) / Ag
ref#1 // external analyte solution / Eb=E1-E2 / ref#2
E1
E2
H+ gradient across the glass; Na+ is the charge carrier at the internal dry part of the membrane
soln glass soln glass
H+ + Na+Gl-  Na+ + H+Gl-

## Slide 9: pH elecctrode glass
Corning 015 is 22% Na2O, 6% CaO, 72% SiO2
Glass must be hygroscopic – hydration of the glass is critical for pH function
The glass surface is predominantly H+Gl- (H+ on the glass) and the internal charge is carried by Na+
glass
H+Gl-
H+Gl-
H+Gl-
H+Gl-
H+Gl-
H+Gl-
H+Gl-
H+Gl-
Na+Gl-
Na+Gl-
E1
E2
Analyte solution
Reference solution

## Slide 10
Values of NIST primary-standard pH solutions from 0 to 60 oC
pH = - log {H+}
K = reference and junction potentials

## Slide 11: Ion Specific Electrodes (ISE’s)
Most utilize a membrane which selects for specific ion(s) – H+, Ca2+, K+, S2-, F-, etc.
This is done through either ion exchange, crystallization, or complexation of the analyte with the electrode surface
Instead of measuring the potential of the galvanic cell, this relates more to a type of junction potential due to separation of an ion

## Slide 12: Electrochemistry - electron transfer reactions
1) Chemical Reactions
4 Fe2+ + O2 + 4 H+  4 Fe3+ + 2 H2O
2) Electrochemical cells - composed of oxidation and reduction half reactions
Fe2+  Fe3+ + e- ANODIC
O2 + 4 H+ + 4 e-  2 H2O CATHODIC
a) Galvanic (Voltaic) cell - thermodynamically favorable or spontaneous (G < 0)
e.g., batteries, pH and ion selective electrode (ISE) measurements
b) Electrolytic cell - non-spontaneous or thermodynamically unfavorable reactions (G > 0) are made to occur with batteries (EAPPL = E applied)
e.g., electrolysis, electroplating, voltammetry

## Slide 13: Electrolytic Cell
Forcing a redox reaction to go in a particular direction by APPLYING the energy required to go forward
Fe2+ + 2e-  Fe0
APPLIED
Potential!
Reaction would not normally go in water

## Slide 14: Environmental Voltammetry
Apply a potential and measure a current response (a half-reaction, involving an e-)
Specific species react at/with the electrode surface AT a specific potential
That potential is from either the equilibrium of a half reaction or a kinetic effect where reaction requires some extra energy to proceed (called overpotential)

## Slide 15
Three Electrode System
Potentiostat
Working Electrode
Reference Electrode
Counter Electrode
Counter Electrode – Pt, facilitates e- flow
Reference Electrode –Ag/AgCl ‘anchors’ V
Working Electrode – Au(Hg) amalgam, surface where specific reactions occur

## Slide 16
V
time
V
time
V
time
Cyclic Voltammetry
Stripping Voltammetry
Square-Wave Voltammetry
Initial ‘plating’ reaction:
Hg + HS-  HgS + H+ + 2 e-
For x seconds followed by reaction
HgS + 2e-+ H+  Hg + HS-
‘Normal’ S-shape curve or peak (due to other processes)
Forward reaction followed by back reaction with V drop  voltammogram dependent on ‘pulse’ height

## Slide 17: Spectroscopy
Exactly how energy is absorbed and reflected, transmitted, or refracted changes the info and is determined by different techniques
sample
Reflected
spectroscopy
Transmittance
spectroscopy
Raman
Spectroscopy

## Slide 18: Light Source
Light shining on a sample can come from different places (in lab from a light, on a plane from a laser array, or from earth shining on Mars from a big laser)
Can ‘tune’ these to any
wavelength or range of
wavelengths
IR image of Mars
Olivine is purple

## Slide 19: Spectroscopy
Beer’s Law:
A = e l c
Where Absorbance, A, is equal to the product of the path length, concentration, c, and molar absorptivity, e

## Slide 20: Causes of Absorption
Molecular or atomic orbitals absorb light, kicks e- from stable to excited state
Charge transfer or radiation (color centers)
Vibrational processes – a bond vibrates at a specific frequency  only specific bonds can do absorb IR though (IR active)

## Slide 21: Emission Spectroscopy
Measurement of the energy emitted upon relaxation of an excited state to a lower state (can be the ground state)
How to generate an excitation – shoot it with high energy particles – UV, X-rays, or heat it in flame or plasma

## Slide 22: Inductively Coupled Plasma
Introduction of molecules in a plasma creates excitations and emits light in the UV and Visible ranges that correspond to elements
Plasma is 7000 degrees – molecules get broken up, the individual elements create the light emission

## Slide 23: Raman Spectroscopy
Another kind of spectroscopy which looks at a scattering effect and what that tells us about the chemistry, oxidation state, and relative proportions of different ions

## Slide 24: Nuclear Magnetic Resonance Spectroscopy (NMR)
NMR is useful for determining short-range cation ordering in minerals.
The NMR spectrometer can be tuned to examine the nucleus of mineralogical interest (e.g. aluminosilicates (27Al, 29Si, 23Na), oxides (17O, 25Mg, etc.), phosphates (31P), hydrous minerals (1H, 19F)).
NMR is particularly useful for cations that can not be distinguished by X-ray methods, such as Si/Al ordering in aluminosilicates

## Slide 25: XANES and EXAFS
X-ray adsorption near-edge spectroscopy and Extended X-ray adsorption Fine Structure, commonly done with synchrotron radiation because the higher energy X-ray yields more precise data
X-ray techniques which look at the fine details of X-ray interactions with minerals
Sensitive to oxidation states and specific bonding environments

## Slide 27: Chromatography
Analyte separation as it moves through a material followed by analysis (spectroscopic or electrochemical)

## Slide 28: Separation
Interaction of analyte with stationary phase based on charge density, hydrophobicity, or size
Analyte displaced across stationary phase by an eluent
Eluent can be ionic (HCO3-), organic (Methanol), gas (Helium)

## Slide 29: chromatograph
Peak separation a function of analyte, stationary phase, eluent composition and flow rate
Goal is to maximize peak separation